## Slide 1
Group 18
Joey Sleppy
John Foster
Zeke rosenbluth
RF Energy Harvesting
for Medical Devices
Notes: Joe

## Slide 2: The Problem is… Batteries Die
Notes: Joe

## Slide 3: Motivation
Notes: Joe

## Slide 4: Motivation
Reduce stress for user’s who rely on batteries for their health
Offer a potential power source to next-gen medical devices such as bio-sensors and artificial organs
Industrial market trends suggesting this could be worth a lot of money
Notes: Joe

## Slide 5: Goals & Objectives
Consistent Communication with user
Little to no user input required
Invisible ~ Comfortable ~ Lightweight

## Slide 6: Specifications
Component | Parameter | Spec
Circuit Size (Wearable) | Size | <4in^2
Output Power | Magnitude & Efficiency | >1mW 25% System Efficiency
Monitoring | Accuracy | 95% True Value
Monitoring | Frequency | Upon user request
Development | Cost | $1,500
Notes: Joe

## Slide 7: A Tale of Two Loads
Larger Load (Insulin Pump)
Requires more support systems
Power monitoring flaw
Small Loads (Sensors)
Require voltage & little current
Easier to scale to broader impacts
Notes: Joe

## Slide 8: NFC Communication Basics
Little to no power consumption
Simple mobile phone integration
No communication overhead
Notes: Cam

## Slide 9
1) RF Signal Sources
2) Signals Received to Antenna
3) Energy Harvested & Delivered
Powercast IC
4a) Data read from sensors
NFC to App
Block Diagram
5) Data Transmission
Wireless Connection
Wired Connection
Micro-controller
Thermal Sensor
Heartrate Sensor
Notes: Joe

## Slide 10
1) RF Signal Sources
2) Signals Received to Antenna
3) Energy Harvested & Delivered
Powercast IC
4a) Data read from sensors
NFC to App
Block Diagram
5) Data Transmission
Wireless Connection
Wired Connection
Micro-controller
Thermal Sensor
Heartrate Sensor
Notes: Joe

## Slide 11: RF Signals
Signal Source | Avg. Strength | Abundance (%) | Volatility | Score
Wi-Fi | -50dBm | 50% | 90% | -22.5
Cellular Receiving | -90dBm | 95% | 90% | -76.95
Cellular Transmitting | 0dBm | 100% | 100% | 0
Score = strength * abundance * volatility = Higher is Better
Radio signals are weak
TV signals are tough to find
Notes: Joe

## Slide 12: Antenna Basics
Parameter:
Monopole
Medium Gain
Small Size
High Efficiency
Notes: Joe

## Slide 13: Antenna Selection
Specifications:
Chip
<70% Efficiency
2dBi Gain
100mm^2
W3012 Ceramic Chip
by Pulse Electronics
Notes: Joe

## Slide 14
Most Efficient
Around 900MHz
Return Loss
of -6dB is bad
Notes: Joe

## Slide 15: Actual Antenna Selection
Specifications:
Whip
  Omni-directional
3dB Gain
3.22” in Length
11.96dB Return Loss
  <10% of signal returned
A1101R09C Whip
by Anaren
Notes: Joe – needs updating here

## Slide 16
1) RF Signal Sources
2) Signals Received to Antenna
3) Energy Harvested & Delivered
Powercast IC
4a) Data read from sensors
NFC to App
Block Diagram
5) Data Transmission
Wireless Connection
Wired Connection
Micro-controller
Thermal Sensor
Heartrate Sensor
Notes: ca,m – be sure to mention that power cast output is enough to power the microcontroller and two sensors and that they are connected in parallel requiring voltge and little current. That the microcontroller needs a buck first.

## Slide 17: Powercast IC
2110B Receiver By Powercast Co.
Notes: cam

## Slide 18: Selection: Powercast IC
2110B Receiver By Powercast Co.
Optimized for Cellular Signal
Input Power < Cellular Signal Power
Easily Controlled output voltage with Resistor
Notes: cam

## Slide 19: Powercast IC
2110B Receiver By Powercast Co.
C= 15*Vo*Io*ton
Rout = 1M(Vo -1.21)
(3.32-Vo)
Notes: Cam – mention that this cap should actually be huge and our learning from it – point out the huge cap. They will like seeing our problem solved skills

## Slide 20
1) RF Signal Sources
2) Signals Received to Antenna
3) Energy Harvested & Delivered
Powercast IC
4a) Data read from sensors
NFC to App
Block Diagram
5) Data Transmission
Wireless Connection
Wired Connection
Micro-controller
Thermal Sensor
Heartrate Sensor
Notes: Cam

## Slide 21: Sensors
Temp Sensor:
Operates between 2.3V and 5.5V
Output data to ADC linearly in microcontroller
Notes: Cam update this slide please – pics, specs, etc…

## Slide 22: Sensors
Notes: Cam update this slide please – pics, specs, etc…

## Slide 23: Microcontroller
Specifications RF430FRL152h:
Entirely powered from host NFC tag
On board MSP430 with ADC to easily read sensor output data
ISO15693 Protocol enbaled
Notes: Cam

## Slide 24
Notes: Cam

## Slide 25
1) RF Signal Sources
2) Signals Received to Antenna
3) Energy Harvested & Delivered
Powercast IC
4a) Data read from sensors
NFC to App
Block Diagram
5) Data Transmission
Wireless Connection
Wired Connection
Micro-controller
Thermal Sensor
Heartrate Sensor
Notes: Zeek

## Slide 26: Software
Internal Sensing System
  RF430FRL152H
External Sensing System
  Android Application
Notes: Zeek

## Slide 27: Software: Sensing System
Programming done in Code Composer Studio
RF430FRL152H
  2 pages of memory designated for RF messages
    0xF868 – 0xFC33, 0xFC34 – 0xFFFF
Notes: Zeek

## Slide 28: Software: Sensing System
Input:
  Voltage level (ADC)
System:
  Low power mode
  Using firmware of RF430
    ISO15693 Protocol
Notes: Zeek

## Slide 29: ISO15693
Notes: Zeek

## Slide 30: Storage
Reads values from sensor until buffer is full
Phone sends read command for each block
Phone sends write command to config block
Notes: Zeek

## Slide 31: Software: Sensing System
Programmed in Android Studio
Take data from RF430
  Plot values on a line graph
  Display average temperature recorded
Notes: Zeek

## Slide 32: Software: Android App
AndroidManifest.xml
  Send NFC action tag discovered to Main Activity
  Enable NFC permissions
MainActivity
  Send ISO15693 commands and display values
Notes: Zeek

## Slide 33: Schematic
RF Energy Harvesting
Power Monitoring
User Communication
Cell signals
received
& harvested
Voltage dropped from 3.4V to 1.5V (buck)
Heartrate Sensor
Temperature
Sensor
Data read by microcontroller
Notes: Cam

## Slide 34
Dimensions:
1.2” by 1.8”
Notes: Cam

## Slide 35: Prototyping
Goals:
Prove concept
Generate ~3V
Determine C2
Expected Behavior & Performance
Antenna
Powercast
Timing
Capacitor
Notes: Cam

## Slide 36: Prototyping Test Results
Results
~3mW peak
<1% of source harvested
Notes: Cam

## Slide 37: Administration
Notes: Joe

## Slide 38: Work Distribution
Joe Sleppy
PCB Design
Energy Harvesting
Zeke Rosenbluth
Monitoring
NFC
John Foster
Energy Storage
Hardware Software Integration
HARDWARE SOFTWARE
Notes: Joe

## Slide 39: Budget
Component | Cost per unit | Total Cost
Components | $800 | $700
PCB Fabrication | $700 (first attempt) $60 (second attempt) | $760
Equipment | $100 | $200
Development Board & Transmitter | $150 | $150
Misc. | $300 | $300
Total Cost |  | $2,110
Notes: Joe need updating

## Slide 40: Progress
Chart: Project Progress
| Research | Design | Prototyping | Optimization | Total
Series 1 | 100 | 95 | 100 | 75 | 92.5
Notes: joe

## Slide 41: Future Work Section
Antenna selection and circuit implementation improvement & optimization
Reduction in size
Investigation of sound waves and other alternative sources vs RF signals
Notes: joe

## Slide 42: Questions?
Notes: joe