## Slide 1: CFS Chlorophyll Fluorescence Spectrometer
Group 1
David Maria, Computer Engineering
Luke Preston, Electrical Engineering
Robert Bernson, Photonic Science & Engineering
Samuel Knight, Photonic Science & Engineering

## Slide 2: Motivation
Plants are vital to human culture, so they must be kept healthy
Hard to tell if a plant is sick, especially if it has green leaves
Check chlorophyll for photosynthetic performance
Traditional lab tests take too long, chemical-based
CFS is non-chemical, faster than lab test
Easy for non-technical user to understand results

## Slide 3: Goals & Objectives
Deliver a low-cost alternative to commercial market products
Make a device which is easy to transport, use, and store (portable)
Simple, modest design that a casual user with a smartphone can use
Robust quality to sustain normal wear-and-tear
  Solid-state components to prevent expensive repairs
Mobile application for casual user to observe results from an app
  Analysis includes optical spectrum and classification, i.e., “Good health”, “Fair health”, “Poor health”, etc.

## Slide 4: Specifications
Parameter | Metric | Reason
Manufacturing cost | ≤ 500 USD | Meets cost goal
Size/Volume | ≤ 4000 cm3 | Meets size goal
Power delivery | ≤ 5 W | Meets heat production
Wireless communication | Bluetooth 4.0 Bluetooth Low Energy mechanism | Facilitates communication to smartphone app
Radio power consumption | ≤ 50 mW | Low power consumption for power circuit
Sample analysis time | ≤ 60 seconds | Fast results
Spectrum analyzed | 600 – 700 nm | Chlorophyll's peak emission spectrum
Notes: David: Analysis time, wireless comms Luke: Power consumption, battery life, power delivery Bobby: Cost, volume, spectrum

## Slide 5
Block Diagram

## Slide 6: Optics

## Slide 7: Laser Housing Section
Houses laser module and cuvette containing sample
Laser center λ = 405nm
Separated from spectroscopy section by optical slit
Two mounts
  One for laser module
  One for 10 mm quartz cuvette

## Slide 8: Chlorophyll Sample Preparation
Process:
  Pick a leaf from a known species of plant
  Break up the leaf into smaller pieces and weigh out a 1-gram sample
  Add the sample to a mixture containing 150mL of acetone
  Muddle the mixture together for 1 minute and pour through a coffee filter
  The remaining solution is added to a cuvette to be tested by the CFS

## Slide 9: Laser Module
Purpose: Provide light source for fluorescence
Manufacturer | Zhuhai MZLASER Tech Co LTD
Part No. | MZH8340550D-AL01A
Cost | $9.00
Wavelength | 405 nm
Beam Size | 8 mm @ 10 m
Power | 50mW @ 3 V
Vendor | Alibaba
Manufacturer | Wavelength | Cost
Frankfurt Laser Co. | 415 nm | $5000.00
Power Technology Inc. | 420 nm | $6310.00
RPMC | 430 nm | $500.00
Zhuhai MZLASER | 405 nm | $9.00

## Slide 10: Quartz Cuvette
Purpose:
Contain Chlorophyll a Sample
Quartz has a high transmission across the necessary range of wavelengths
Allows 405 nm laser light to pass through and illuminate sample
Allows 600-700 nm fluorescent light to emit through container
Manufacturer | Science Outlet
Part No. | 100-40-10
Cost | $9.99
Volume | 3.5 mL
Dimensions | 12.5 x 12.5 x 45 mm
Thickness | 1.25 mm
Vendor | Amazon
Manufacturer | Volume | Dimensions | Cost
Lichen Cottage | 3.5 mL | 12.5 x 12.5 x 45 mm | $19.94
Alpha Nanotech | 3.5 mL | 12.5 x 12.5 x 45 mm | $21.00
Bitomic | 3.5 mL | 12.5 x 12.5 x 45 mm | $17.50
Science Outlet | 3.5 mL | 12.5 x 12.5 x 45 mm | $9.99
Notes: Quartz has a relatively high transmission at 405nm

## Slide 11: Optical Slit
Purpose:
Allow fluorescent light to pass onto collimating mirror while preventing directional 405 nm laser light from passing through
Ideal thickness: 350 µm
System uses 200 µm since 350 µm would have to be
machined by hand
Image of slit propagates through optical cavity

## Slide 12: Original Spectroscopy Section
Intense angles between optics lead to unexpected astigmatism, clipping at grating interface
Image of slit at the sensor was spread out into a wide rectangle; no spectrum imageable

## Slide 13: New Spectroscopy Section
Grating moved position; angle of incidence changed
Collimating mirror angle decreased
Focusing mirror unchanged
Sensor moved to astigmatized focal point
Light now focuses properly
  No clipping
  No visible astigmatism

## Slide 14: Collimating & Focusing Mirrors
Manufacturer | ThorLabs
Part No. | CM254-050-P01
Cost | $60.34
Scratch-Dig | 40-20
Diameter | 25.4 mm
Focal Length | 50 mm
Vendor | ThorLabs
Purpose:
Collimate light from the pinhole and direct it to the diffraction grating
Focus light from the diffraction grating onto the sensor
Silver-coated, offering >97.5% reflectance from 450 to 2000 nm
For applications near 780 nm, silver exceeds aluminum in quality
F# = 1.98, which captures a lot of light but doesn’t provide much depth of field; difficult to focus
Manufacturer | EFL | Diameter | Cost
Edmund Optics | 50 mm | 25.4 mm | $165.00
Newport | 50 mm | 25.4 mm | $141.00
ThorLabs | 50 mm | 25.4 mm | $60.34

## Slide 15: Diffraction Grating
Manufacturer | MKS ¦ Newport
Part No. | 33066FL01-270R
Cost | $89.00
Grooves/mm | 300
Spectral range | 300 – 1000 nm
Wavelength of maximum efficiency | 500 nm
Size | 12.5 x 25 mm
Vendor | MKS ¦ Newport
Purpose:
Split the light coming from the collimating mirror into its component wavelengths
Grating has an efficiency of 62 – 73% in the CFS’s waveband (600 to 700 nm)
Cheaper than 25 x 25mm grating
Manufacturer | Grv/mm | Size | Cost
ThorLabs | 300 | 25 x 25 mm | $117
Newport | 300 | 25 x 25 mm | $109
Newport | 300 | 12.5 x 25 mm | $89

## Slide 16: CMOS Sensor Array
Manufacturer | ON Semiconductor
Part No. | AR0130CSSC00SPBA0-DR
Cost | $8.38 / unit
Active area | 4.86 x 3.66 mm
Supply Voltage | 1.8 V & 2.8 V
Operating Wavelength | 300 – 950 nm
ADC Depth | 12 bits
Vendor | DigiKey
Purpose:
Observe the light coming from the focusing mirror and transmit pixel intensity profile data to the CPU for processing
History of successes at UCF using CMOS over linear
Operates in analog, digital, or I/O mode
Built-in ADC
68 – 78% QE in the CFS’s waveband (600 – 700 nm)
Monochromatic pixels; only one QE/sensitivity curve
3.66mm
Manufacturer | Sensor Type | Supply Voltage | ADC Depth | Cost
Melexis | Linear | 3 – 3.3 V | 8 bits | $13.65
Renesas | CMOS | 2.25 – 3.63 V | 8 or 12 bits | $20.00
AMS | Linear | 3 – 5 V | 8 bits | $9.80
ON Semi | CMOS | 1.8 V & 2.8 V | 12 bits | $8.38

## Slide 17: MCU

## Slide 18: STM32 ARM CPU
Control laser module operation
Control sensor operation
Control data flow
Low Cost: $8.82
Affordable development board: $8.79
Low Power Consumption: 58.5 mW at 1.95 V
Modern and capable CPU
Supports UART, I2C, SPI
Sufficient GPIO connectivity
CPU | Chip Cost | Dev Board Cost | GPIO
MSP 430FR2672 | $3.85 | $15 | Not Enough
STM32F407VET | $4.41 | $9 | Enough
FS32K144 | $5.00 | $65 | Enough

## Slide 19: Mobile Application

## Slide 20: Bluetooth Module
Provide communication between CFS and smartphone
Bluetooth preferable to an on-board display
Low Power Consumption: Peak power draw of 31.2 mW
Low Voltage: 1.95 Volt supply needed
Low cost: $15.40
Supports Bluetooth Low Energy
Module | Manufacturer | Size | Cost | Power Consumption | Bluetooth Version | Vendor Used
HC-05 | ITead Studio | 12.7 x 27 mm | <$10 | Medium | V2.0 | N/A
RN4020 | Microchip | 11.5 x 19.5 x 2.5 mm | $15 | Very Low | V4.1, BLE | Amazon
ESP32 | Espressif Systems​ | 25.5 × 18 × 3.1 mm | <$20 | Low | V4.2, BLE+EDR | N/A

## Slide 21: Software
Design Choice | Final Selection
Supported Mobile Operating System(s) | Android
Earliest Supported Android API | API level 21
Mobile Application Local Database Technology | SQlite3
Mobile Application Programming Language | Java
Embedded Software Language | C
Embedded Software IDE | System Workbench for STM32

## Slide 22: Embedded Software Logic

## Slide 23: Application Logic

## Slide 24: Application Use Case Diagram

## Slide 25: Entity Relationship Diagram

## Slide 26: Mobile phone GUI
Searching for devices
Analysis Options
Choosing Stored Analyses
Viewing Analysis Results

## Slide 27: Power System

## Slide 28: Power Supply Circuit
Battery & charging circuit
3.3 V for CPU & BT
3.0 V for laser module
1.8 V for CMOS sensor digital power
2.8 V for CMOS sensor analog power

## Slide 29: Printed Circuit Board Design
Schematic separated into 2 PCBs
Sensor mounted in place
Connected by ribbon cable
USB charging
Debugging pinouts added

## Slide 30: Housing Design
Cost to build: $0.00
Critical that the spectroscopy section and pump section be isolated from stray light
Housing never completed and fully enclosed due to COVID-19 shutdown
  Device can only be operated in a completely dark room

## Slide 31: Component & Integration Testing

## Slide 32: Software Development Prototype
STM32F407 Blackboard Development Board
RN4020 Bluetooth Low Energy Module
  UART
OV7670 Image Sensor
  Digital Camera Memory Interface
Android Application
  Developed and Debugged using Android Studio
Image Sensor
Bluetooth Sensor
MCU

## Slide 33: Prototype Used for Testing
(From left to right, top to bottom)
Laser test source
Power supply
Laser pump source
Chlorophyll sample
Collimating mirror
Reflective diffraction grating
USB enabled sensor
Focusing mirror

## Slide 34: Optics Testing
Grating held in front of entrance slit to show modal diffraction
Grating diffracting laser light in multiple modes with mirror pictured (bottom)
Laser (right side) propagating through optical system to screen (top center)

## Slide 35: Initial Integration Testing
Connected AR0130CS sensor to laptop via USB cable
Used built-in Camera program in Windows to view spectrum
Could see fluorescence spectrum from chlorophyll, green light scattering from laser due to sensor not completely isolated

## Slide 36: Fluorescence Calibration
Calibrated using labeled green and red laser sources from a laser light show module
Once peaks were located by pixel, x-axis was shifted and relabeled so peaks were at correct wavelengths, not pixels
Resolution is 5.4 px/nm or 0.185 nm/px

## Slide 37: Fluorescence Testing
Newest sample on top shows peak wavelength 667 nm
  Scattering from laser
Old sample on bottom shows peak wavelength at 665 nm
Chlorophyll fluorescence spectrum peaks range from 662 to 669 nm depending on sample age
Samples are within this range as expected

## Slide 38: Final Integration Testing
Connected calibrated Optics system to Mobile Application
  Bluetooth Low Energy Connection
Chlorophyll Sample
Red Laser
Green Laser

## Slide 39: Current Health Metric
The CFS measures and displays the fluorescence spectrum of Chlorophyll a
  Allows us to determine the peak emission wavelength
  Our device then compares the measured spectrum to an already known healthy spectrum of the same plant species (reference sample)
  If the sample has a peak intensity less than 60% of our reference sample our device has determined it to be unhealthy
Spectrum of an unhealthy sample (black) in comparison to a reference sample (green) both recorded with our device
Notes: Our device determines relative plant health
A measured sample must be compared to a reference sample of the same species which has already been determined to be healthy
Our device determines a measured sample to be unhealthy if its peak intensity is less than 60 % of the peak intensity of the reference sample

## Slide 40: Possible Future Health Metrics
Since we can determine our peak emission wavelength and our peak fluorescence intensity FM, if we were to add more sources to our device it would be possible to determine other chlorophyll fluorescence parameters
  By using a saturating white light it is possible to determine the fluorescence decrease ratio RFd=(FM/FS) – 1
    FS is the steady state fluorescence
    If the RFd > 2.5, sample is getting enough sunlight
  By using a pulsed red LED in conjunction with a saturating white light, it is possible to determine the maximum quantum yield FV/FM
    FV is the variable fluorescence
    0.74 < FV/FM < 0.85, sample is normal (i.e. healthy)
Notes: By knowing our peak emission wavelength, it is possible to find other common chlorophyll fluorescence parameters
However, these parameters would require other light sources not included in our device
Unfortunately, timing and COVID-19 concerns prevented us from adding these sources to our device
Using a white light source we could determine the steady state fluorescence which would allow us to determine the fluorescence decrease ratio
An Rfd value greater than 2.5 tells us the measured sample is getting enough sunlight
Using a pulsed red LED we could determine the variable fluorescence which would allow us to determine the maximum quantum yield
If the measured sample has a Fv/Fm between the values of 0.74 and 0.85 then it is determined to be healthy

## Slide 41: Final Product Specifications
Parameter | Goal Metric | Final Metric | Meet Spec?
Manufacturing cost | ≤ 500 USD | 400 USD | Yes
Size/Volume | ≤ 4000 cm3 | 1958.4 cm3 (19.2 x 12 x 8.5 cm) | Yes
Power delivery | ≤ 5 W | 3.61 W | Yes
Wireless communication | Bluetooth 4.0 Bluetooth Low Energy mechanism | Bluetooth 4.0 Bluetooth Low Energy mechanism | Yes
Radio power consumption | ≤ 50 mW | 39.6 mW | Yes
Sample analysis time | ≤ 60 seconds | 40 seconds | Yes
Spectrum analyzed | 600 – 700 nm | 500 – 740 nm | Yes

## Slide 42: Administrative Content

## Slide 43: Work Distribution
Task | Primary | Secondary | Workload | Major
Pump source section design | Robert | Samuel | Medium | PSE
Chlorophyll sample holder design | Robert | Samuel | Medium | PSE
Spectroscopy section design | Samuel | Robert, Luke | High | PSE
Sensor testing | David, Luke | Robert, Samuel | High | CE, EE, PSE
Integration testing | Samuel, David | Robert | High | PSE, CE
CFS housing design | Samuel | Robert | Medium | PSE
PCB testing and design | Luke | David | High | EE, CE
Power circuit design | Luke | David | Low | EE
Embedded software design | David | Luke | High | CE
Mobile application design | David | Luke | Medium | CE

## Slide 44: Budget & Financials
Component | Supplier | Device Section | Quantity | Total Cost
Laser source | Zhuhai MZLAZER Tech Co LTD. | Laser housing | 2 | $40.38
Cuvette | Scientific Lab Solution | Laser housing | 2 | $20.00
Mirrors | ThorLabs | Optical housing | 2 | $120.68
Diffraction grating | Newport Corporation | Optical housing | 1 | $89.00
CMOS Sensor | ON Semiconductor | Optical housing | 4 | $34.68
PCB | Oshpark LLC | PCB | 3 | $46.00
CPU | STMicroelectronics | PCB | 2 | $8.82
Development board | STMicroelectronics | Dev board | 1 | $8.79
Bluetooth module | Microchip | PCB | 1 | $29.85
Final Cost |  |  |  | $398.20

## Slide 45: Progress
Laser and cuvette not mounted
Diffraction grating not mounted properly
Sensor not mounted properly
Final machining not completed
Device not isolated from light, can only operate in a completely dark room
PCB hardware redesign not received
Chart: Progress Chart
| Total | Testing | Prototype | Design | Research
Series 1 | 0.96 | 1 | 0.85 | 1 | 1

## Slide 46: Q&A